## Slide 1
Wavewin Sniffer 24 Configuration & Polling Software
(Wavewin Sniffer Operations Guide)

## Slide 2
Operations Guide Content Capturing Data

## Slide 3
Capturing Data
Run the Wavewin Sniffer Software
Double Click “Wavewin Sniffer”.

## Slide 4
Capturing Data Continuous Data Capture &
Trigger Data Capture
Place the sensor in
the intended environment

## Slide 5
Step 1: Enter “Channel Title”.
Step 2: Select “Sensor Type”.

## Slide 6
Step 3: Click “File Properties” & check “Save Continuous”. Then click “OK”.

## Slide 7
Step 4: To begin polling click “Start Recorder”.

## Slide 8
Triggering Wavewin Sniffer has 6 different trigger settings:
Trigger Level,
Duration,
Upper Hysteresis,
Lower Hysteresis,
Absolute,
Operator:
greater than (>), less than (<), equal to (=),
not equal to (<>).

## Slide 9
Triggering
Logic | Trigger | Reset
= | Region 2 | Region 1 or Region 3
<> | Region 1 or Region 3 | Region 2
> | Region 3 | Region 1
< | Region 1 | Region 3

## Slide 10
Example of Triggering How to configure Wavewin Sniffer to
trigger at 5.0 Amps RMS
with Hysteresis region of 1.0 Amps and
a duration of 2 ms.

## Slide 11
Step 1: Enter “5” in “Trigger Value” field.

## Slide 12
Step 2: From the drop down list, select RMS.

## Slide 13
Step 3: Enter “2” in “Duration” field.

## Slide 14
Step 4: Enter “0.5” in “Upper Hysteresis” & “0.5” in “Lower Hysteresis” fields.

## Slide 15
Step 5: Select “Absolute” checkbox. This field is helpful in case the sensor was mistakenly mounted in the reverse direction.

## Slide 16
Step 6: Select “>” from “Operator” drop down list.

## Slide 17
Step 7: Save the active configuration to disk.

## Slide 18
Step 8: Click “File Properties”.

## Slide 19
Step 9: For trigger data capture, uncheck “Save Continuous”. Then click “OK”.

## Slide 20
Step 10: Click “Start Recorder” to begin polling.

## Slide 21
End of Operations Guide